## N8800 settings
Configuration Parameters (1st)
Cntr (Control and Self-tunning)
Name | Value range | Description | Default | Actual Settings
SP.Fn |  | Basic configuration of setpoint processing | 0 | 0
| 0 | set-point controller can be switched over to external set-point (-> LOGI/ SP.E)
| 8 | standard controller with external offset (SP.E)
C.tYP |  | Calculation of the process value | 0 | 0
| 0 | standard controller (process value = x1)
| 1 | ratio controller (x1/x2)
| 2 | difference (x1 - x2)
| 3 | Maximum value of x1and x2. It is controlled with the bigger value. At sensor failure it is controlled with the remaining actual value.
| 4 | Minimum value of x1and x2. It is controlled with the smaller value. At sensor failure it is controlled with the remaining actual value.
| 5 | Mean value (x1, x2). With sensor error, controlling is continued with the remaining process value.
| 6 | Switching between x1 and x2 (-> LOGI/ I.ChG)
| 7 | O2 function with constant sensor temperature
| 8 | O2 function with measured sensor temperature
C.Fnc |  | Control behaviour (algorithm) | 1 | 1
| 0 | on/off controller or signaller with one output
| 1 | PID controller (2-point and continuous)
| 2 | ∆ / Y / Off, or 2-point controller with partial/full load switch-over
| 3 | 2 x PID (3-point and continuous)
| 4 | 3-point stepping controller
| 5 | 3-point stepping controller with position feedback Yp
| 6 | continuous controller with integrated positioner
C.dif |  | Output action of the PID controller derivative action | 0 | 1
| 0 | Derivative action acts only on the measured value.
| 1 | Derivative action only acts on the control deviation (set-point is also differentiated)
nAn |  | Manual operation permitted | 0 | 1
| 0 | no
| 1 | yes (->LOGI / nAn)
C.Act |  | Method of controller operation | 0 | 1
| 0 | inverse, e.g. heating
| 1 | direct, e.g. cooling
FAIL |  | Behaviour at sensor break | 1 | 1
| 0 | controller outputs switched off
| 1 | y = Y2
| 2 | y = mean output. The maximum permissible output can be adjusted with parameter Ym.H. To prevent determination of inadmissible values, mean value formation is only if the control deviation is lower than parameter L.Ym.
rnG.L | -1999...9999 | X0 (low limit range of control) 1 | -100 | 800
rnG.H | -1999...9999 | X100 (high limit range of control) 1 | 1200 | 1200
CYCL |  | Characteristic for 2-point- and 3-point-controllers | 0 | 0
| 0 | standard
| 1 | water cooling linear
| 2 | water cooling non-linear
| 3 | with constant cycle (see page 51)
tunE |  | Auto-tuning at start-up (see page 16 ) | 0 | 1
| 0 | At start-up with step attempt, at set-point with impulse attempt
| 1 | At start-up and at set-point with impulse attempt. Setting for fast controlled systems (e.g. hot runner control)
| 2 | Always step attempt at start-up
Strt |  | Start of auto-tuning | 0 | 1
| 0 | Manual start of auto-tuning
| 1 | Manual or automatic start of auto-tuning at power on or when oscillating is detected
Adt0 |  | Optimization of T1, T2 (only visible with 8800/8840 configurator!) | 0
| 0 | Automatic optimization
| 1 | No optimization
1 rnG.L and rnG.H are indicating the range of control on which e.g. the self-tuning is refering
InP.1
Name | Value range | Description | Default | Last settings
I.Fnc |  | INP1 function selection | 7 | 7
| 0 | No function (following INP data are skipped)
| 1 | Heating current input
| 2 | External set-point SP.E (switch-over -> LOGI/ SP.E)
| 3 | Position feedback Yp
| 4 | Second process value x2 (ratio, min, max, mean)
| 5 | External positioning value Y.E (switch-over r LOGI / Y.E)
| 6 | No controller input (e.g. limit signalling instead)
| 7 | Process value x1
s.typ |  | Sensor type selection | 1 | 30
| 0 | thermocouple type L (-100...900°C) , Fe-CuNi DIN
| 1 | thermocouple type J (-100...1200°C) , Fe-CuNi
| 2 | thermocouple type K (-100...1350°C), NiCr-Ni
| 3 | thermocouple type N (-100...1300°C), Nicrosil-Nisil
| 4 | thermocouple type S (0...1760°C), PtRh-Pt10%
| 5 | thermocouple type R (0...1760°C), PtRh-Pt13%
| 6 | thermocouple type T (-200...400°C), Cu-CuNi
| 7 | thermocouple type C (0...2315°C), W5%Re-W26%Re
| 8 | thermocouple type D (0...2315°C), W3%Re-W25%Re
| 9 | thermocouple type E (-100...1000°C), NiCr-CuNi
| 10 | thermocouple type B (0/100...1820°C), PtRh-Pt6%
| 18 | special thermocouple
| 20 | Pt100 (-200.0 ... 100,0 °C)
| 21 | Pt100 (-200.0 ... 850,0 °C)
| 22 | Pt1000 (-200.0 ... 850.0 °C)
| 23 | special 0...4500 Ohm (preset to KTY11-6)
| 24 | special 0...450 Ohm
| 30 | 0...20mA / 4...20mA 1
| 40 | 0...10V / 2...10V 1
| 41 | special -2,5...115 mV 1
| 42 | special -25...1150 mV 1
| 50 | potentiometer 0...160 Ohm 1
| 51 | potentiometer 0...450 Ohm 1
| 52 | potentiometer 0...1600 Ohm 1
| 53 | potentiometer 0...4500 Ohm 1
S.Lin |  | Linearization (only at S.tYP = 23 (KTY 11-6), 24 (0...450 ), 30 (0..20mA), 40 (0..10V), 41 (0...100mV) and 42 (special -25...1150 mV)) | 0 | 0
| 0 | none
| 1 | Linearization to specification. Creation of linearization table with 8800/8840 configurator (engineering tool) possible. The characteristic for KTY 11-6 temperature sensors is preset.
Corr |  | Measured value correction / scaling | 0 | 3
| 0 | Without scaling
| 1 | Offset correction (at CAL level)
| 2 | 2-point correction (at CAL level)
| 3 | Scaling (at PArA level)
In.F | -1999...9999 | Alternative value for error at INP1 | OFF | OFF
fAI1 |  | Forcing INP1 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
1 with current and voltage input signals, scaling is required (see chapter 5.3)
InP.2
Name | Value range | Description | Default | Last settings
I.Fnc |  | Function selection of INP2 | 1 | 0
| 0 | no function (subsequent input data are skipped)
| 1 | heating current input
| 2 | External set-point SP.E (switch-over -> LOGI/ SP.E)
| 3 | Position feedback Yp
| 4 | Second process value x2 (ratio, min, max, mean)
| 5 | External positioning value Y.E (switch-over r LOGI / Y.E)
| 6 | no controller input (e.g. transmitter input instead)
| 7 | Process value x1
S.tYP |  | Sensor type selection | 30
| 30 | 0...20mA / 4...20mA 1
| 31 | 0...50mA AC 1
| 50 | Potentiometer ( 0...160 Ohm) 1
| 51 | Potentiometer ( 0...450 Ohm) 1
| 52 | Potentiometer ( 0...1600 Ohm) 1
| 53 | Potentiometer ( 0...4500 Ohm) 1
Corr |  | Measured value correction / scaling | 0
| 0 | Without scaling
| 1 | Offset correction (at CAL level)
| 2 | 2-point correction (at CALlevel)
| 3 | Scaling (at PArA level)
In.F | -1999...9999 | Alternative value for error at INP2 | OFF
fAI2 |  | Forcing INP2 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
1 with current and voltage input signals, scaling is required (see chapter 5.3)
Lin (Limit value functions)
Name | Value range | Description | Default | Last settings
Fnc.1 |  | Function of limit 1 | 1 | 1
| 0 | switched off
| 1 | measured value monitoring
| 2 | Measured value monitoring + alarm status storage. A stored limit value can be reset via error list, è-key, Ò-key or a digital input ( -> LOGI/ Err.r)
| 3 | signal change (change/minute)
| 4 | signal change and storage (change/minute)
Src.1 |  | Source of limit 1 | 1 | 0
| 0 | process value
| 1 | control deviation xw (process value - set-point)
| 2 | control deviation xw (with suppression after start-up and set-point change) After switch-on or set-point changing, the alarm output is suppressed, until the process value is within the limits for the first time. At the latest after elapse of time 10 ti1the alarm is activated (ti1 = integral time 1; parameter r Cntr) If ti1 is switched off (ti1 = 0), this is interpreted as Î , i.e. the alarm is not activated, before the process value was within the limits once.
| 3 | measured value INP1
| 4 | measured value INP2
| 5 | measured value INP3
| 6 | effective setpoint Weff
| 7 | correcting variable y (controller output)
| 8 | control variable deviation xw (actual value - internal setpoint) = deviation alarm to internal setpoint
| 9 | difference x1 - x2 (utilizable e.g. in combination with process value function “mean value” for recognizing aged thermocouples
Fnc.2 |  | Function of limit 2 | 0 | 0
| 0 | switched off
| 1 | measured value monitoring
| 2 | Measured value monitoring + alarm status storage. A stored limit value can be reset via error list, è-key, Ò-key or a digital input ( -> LOGI/ Err.r)
| 3 | signal change (change/minute)
| 4 | signal change and storage (change/minute)
Src.2 |  | Source of limit 2 | 0
| 0 | process value
| 1 | control deviation xw (process value - set-point)
| 2 | control deviation xw (with suppression after start-up and set-point change) After switch-on or set-point changing, the alarm output is suppressed, until the process value is within the limits for the first time. At the latest after elapse of time 10 ti1the alarm is activated (ti1 = integral time 1; parameter r Cntr) If ti1 is switched off (ti1 = 0), this is interpreted as Î , i.e. the alarm is not activated, before the process value was within the limits once.
| 3 | measured value INP1
| 4 | measured value INP2
| 5 | measured value INP3
| 6 | effective setpoint Weff
| 7 | correcting variable y (controller output)
| 8 | control variable deviation xw (actual value - internal setpoint) = deviation alarm to internal setpoint
| 9 | difference x1 - x2 (utilizable e.g. in combination with process value function “mean value” for recognizing aged thermocouples
Fnc.3 |  | Function of limit 3 | 0 | 0
| 0 | switched off
| 1 | measured value monitoring
| 2 | Measured value monitoring + alarm status storage. A stored limit value can be reset via error list, è-key, Ò-key or a digital input ( -> LOGI/ Err.r)
| 3 | signal change (change/minute)
| 4 | signal change and storage (change/minute)
Src.3 |  | Source of limit 3 | 0
| 0 | process value
| 1 | control deviation xw (process value - set-point)
| 2 | control deviation xw (with suppression after start-up and set-point change) After switch-on or set-point changing, the alarm output is suppressed, until the process value is within the limits for the first time. At the latest after elapse of time 10 ti1 the alarm is activated (ti1 = integral time 1; parameter r Cntr) If ti1 is switched off (ti1 = 0), this is interpreted as Î , i.e. the alarm is not activated, before the process value was within the limits once.
| 3 | measured value INP1
| 4 | measured value INP2
| 5 | measured value INP3
| 6 | effective setpoint Weff
| 7 | correcting variable y (controller output)
| 8 | control variable deviation xw (actual value - internal setpoint) = deviation alarm to internal setpoint
| 9 | difference x1 - x2 (utilizable e.g. in combination with process value function “mean value” for recognizing aged thermocouples
HC.AL |  | Alarm heat current function (INP2) | 0 | 0
| 0 | switched off
| 1 | Overload short circuit monitoring
| 2 | Break and short circuit monitoring
LP.AL |  | Monitoring of control loop interruption for heating | 0 | 0
| 0 | switched off / inactive
| 1 | active. If ti1=0 LOOP alarm is inactive!
dAc.A |  | DAC alarm function | 0
| 0 | DAC alarm switched off / inactive
| 1 | DAC alarm active
Hour | OFF...999999 | Operating hours (only visible with 8800/8840 configurator!) | OFF
Swit | OFF...999999 | Output switching cycles (only visible with 8800/8840 configurator!) | OFF
Out.1
N+250:270ame | Value range | Description | Default | Last settings
0.Act |  | Method of operation of output OUT1 | 0 | 0
| 0 | direct / normally open
| 1 | inverse / normally closed
Y.1 |  | Controller output Y1 | 1 | 0
| 0 | not active
| 1 | active
Y.2 |  | Controller output Y2 | 0
| 0 | not active
| 1 | active
Lin.1 |  | Limit 1 signal | 0 | 1
| 0 | not active
| 1 | active
Lin.2 |  | Limit 2 signal | 0
| 0 | not active
| 1 | active
Lin.3 |  | Limit 3 signal | 0
| 0 | not active
| 1 | active
dAc.A |  | Valve monitoring (DAC) | 0
| 0 | not active
| 1 | active
LP.AL |  | Interruption alarm signal (LOOP) | 0
| 0 | not active
| 1 | active
HC.AL |  | Heat current alarm signal | 0
| 0 | not active
| 1 | active
HC.SC |  | Solid state relay (SSR) short circuit signal | 0
| 0 | not active
| 1 | active
FAi.1 |  | INP1 error signal | 0 | 0
| 0 | not active
| 1 | active
FAi.2 |  | INP2 error signal | 0
| 0 | not active
| 1 | active
FAi.3 |  | INP3 error signal | 0
| 0 | not active
| 1 | aktiv
fOut |  | Forcing OUT1 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
Out.2 Configuration parameters Out.2 = Out.1 except for: Default Y.1 =0 Y.2 =1
N+250:270ame | Value range | Description | Default | Last settigs
0.Act |  | Method of operation of output OUT2 | 0 | 0
| 0 | direct / normally open
| 1 | inverse / normally closed
Y.1 |  | Controller output Y1 | 0 | 0
| 0 | not active
| 1 | active
Y.2 |  | Controller output Y2 | 1
| 0 | not active
| 1 | active
Lin.1 |  | Limit 1 signal | 0 | 0
| 0 | not active
| 1 | active
Lin.2 |  | Limit 2 signal | 0
| 0 | not active
| 1 | active
Lin.3 |  | Limit 3 signal | 0
| 0 | not active
| 1 | active
dAc.A |  | Valve monitoring (DAC) | 0
| 0 | not active
| 1 | active
LP.AL |  | Interruption alarm signal (LOOP) | 0
| 0 | not active
| 1 | active
HC.AL |  | Heat current alarm signal | 0
| 0 | not active
| 1 | active
HC.SC |  | Solid state relay (SSR) short circuit signal | 0
| 0 | not active
| 1 | active
FAi.1 |  | INP1 error signal | 0 | 0
| 0 | not active
| 1 | active
FAi.2 |  | INP2 error signal | 0
| 0 | not active
| 1 | active
FAi.3 |  | INP3 error signal | 0
| 0 | not active
| 1 | aktiv
fOut |  | Forcing OUT2 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
Out.3
Name | Value range | Description | Default | Last settings
O.tYP |  | Signal type selection OUT3 | 0 | 5
| 0 | relay / logic (only visible with current/logic voltage)
| 1 | 0 ... 20 mA continuous (only visible with current/logic/voltage)
| 2 | 4 ... 20 mA continuous (only visible with current/logic/voltage)
| 3 | 0...10 V continuous (only visible with current/logic/voltage)
| 4 | 2...10 V continuous (only visible with current/logic/voltage)
| 5 | transmitter supply (only visible without OPTION)
0.Act |  | Method of operation of output OUT3 (only visible when O.TYP=0) | 1
| 0 | direct / normally open
| 1 | inverse / normally closed
Y.1 |  | Controller output Y1 (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Y.2 |  | Controller output Y2 (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Lin.1 |  | Limit 1 signal (only visible when O.TYP=0) | 1
| 0 | not active
| 1 | active
Lin.2 |  | Limit 2 signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Lin.3 |  | Limit 3 signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
dAc.A |  | Valve monitoring (DAC) (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
inp.1 |  | Interruption alarm signal (LOOP) (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
HC.AL |  | Heating current alarm signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
HC.SC |  | Solid state relay (SSR) short circuit signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
FAi.1 |  | INP1 error (only visible when O.TYP=0) | 1
| 0 | not active
| 1 | active
FAi.2 |  | INP2 error (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
FAi.3 |  | INP3 error (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | aktiv
Out.0 | -1999...9999 | Scaling of the analog output for 0% (0/4mA or 0/2V, only visible when O.TYP=1..5) | 0
Out.1 | -1999...9999 | Scaling of the analog output for 100% (20mA or 10V, only visible when O.TYP=1..5) | 100
O.Src |  | Signal source of the analog output OUT3 (only visible when O.TYP=1..5) | 1
| 0 | not used
| 1 | controller output y1 (continuous)
| 2 | controller output y2 (continuous)
| 3 | process value
| 4 | effective set-point Weff
| 5 | control deviation xw (process value - set-point)
| 6 | measured value position feedback Yp
| 7 | measured value INP1
| 8 | measured value INP2
| 9 | measured value INP3
fOut |  | Forcing OUT3 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
Out.4
Name | Value range | Description | Default | Last settings
O.tYP |  | Signal type selection OUT4 | 0 | 2
| 0 | relay / logic (only visible with current/logic voltage)
| 1 | 0 ... 20 mA continuous (only visible with current/logic/voltage)
| 2 | 4 ... 20 mA continuous (only visible with current/logic/voltage)
| 3 | 0...10 V continuous (only visible with current/logic/voltage)
| 4 | 2...10 V continuous (only visible with current/logic/voltage)
| 5 | transmitter supply (only visible without OPTION)
O.Act |  | Method of operation of output OUT4 (only visible when O.TYP=0) | 0
| 0 | direct / normally open
| 1 | inverse / normally closed
Y.1 |  | Controller output Y1 (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Y.2 |  | Controller output Y2 (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Lin.1 |  | Limit 1 signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Lin.2 |  | Limit 2 signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
Lin.3 |  | Limit 3 signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
dAc.A |  | Valve monitoring (DAC) (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
LP.AL |  | Interruption alarm signal (LOOP) (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
HC.AL |  | Heat current alarm signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
HC.SC |  | Solid state relay (SSR) short circuit signal (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
FAi.1 |  | INP1 error (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
FAi.2 |  | INP2 error (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | active
FAi.3 |  | INP3 error (only visible when O.TYP=0) | 0
| 0 | not active
| 1 | aktiv
Out.0 | -1999...9999 | Scaling of the analog output for 0% (0/4mA or 0/2V, only visible when O.TYP=1..5) | 0 | 4
Out.1 | -1999...9999 | Scaling of the analog output for 100% (20mA or 10V, only visible when O.TYP=1..5) | 100 | 20
O.src |  | Signal source of the analog output OUT4 (only visible when O.TYP=1..5) | 0 | 1
| 0 | not used
| 1 | controller output y1 (continuous)
| 2 | controller output y2 (continuous)
| 3 | process value
| 4 | effective set-point Weff
| 5 | control deviation xw (process value - set-point)
| 6 | measured value position feedback Yp
fOut |  | Forcing OUT1 (only visible with 8800/8840 configurator!) | 0
No forcing
Forcing via serial interface
0.FAI |  |  |  | 0
LOGI (Digital imputs)
Name | Value range | Description | Default | Last settings
L_r |  | Local / Remote switching (Remote: adjusting of all values by front keys is blocked) | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 1 | always active
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
SP.2 |  | Switching to second setpoint SP.2 | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | è - key switches
SP.E |  | Switching to external setpoint SP.E | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 1 | always active
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | è - key switches
Y2 |  | Y/Y2 switching | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | è - key switches
| 6 | Ò - key switches
Y.E |  | Switching to fixed control output Y.E | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 1 | always activated (manual station)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | è - key switches
| 6 | Ò - key switches
nAn |  | Automatic/manual switching | 0 | 6
| 0 | no function (switch-over via interface is possible)
| 1 | always activated (manual station)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
| 6 | Manual - key switches
C.oFF |  | Switching off the controller | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | è - key switches
| 6 | Ò - key switches
n.Loc |  | Blockage of hand function | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
Err.r |  | Reset of all error list entries | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
| 6 | Manual - key switches
Pid.2 |  | Switching of parameter set (Pb, ti, td) | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
I.ChG |  | Switching of the actual process value between Inp1 and X2 | 0 | 0
| 0 | no function (switch-over via interface is possible)
| 2 | DI1 switches
| 3 | DI2 switches (only visible with OPTION)
| 4 | DI3 switches (only visible with OPTION)
| 5 | F - key switches
di.Fn |  | Function of digital inputs (valid for all inputs) | 0 | 0
| 0 | direct
| 1 | inverse
| 2 | toggle key function
fDI1 |  | Forcing di1 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
fDI2 |  | Forcing di2 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
fDI3 |  | Forcing di3 (only visible with 8800/8840 configurator!) | 0
| 0 | No forcing
| 1 | Forcing via serial interface
othr (Display, operation, interface)
Name | Value range | Description | Default | Last settings
bAud |  | Baudrate of the interface (only visible with OPTION) | 2
| 0 | 2400 Baud
| 1 | 4800 Baud
| 2 | 9600 Baud
| 3 | 19200 Baud
Addr | 1...247 | Address on the interace (only visible with OPTION) | 1
PrtY |  | Data parity on the interface (only visible with OPTION) | 1
| 0 | no parity (2 stop bits)
| 1 | even parity
| 2 | odd parity
| 3 | no parity (1 stopbit)
dELY | 0...200 | Delay of response signal [ms] (only visible with OPTION) | 0
O2 |  | Entering parameter for O2 in ppm or % | 0
| 0 | Parameter for O2-function in ppm
| 1 | Parameter for O2-function in %
Unit |  | Unit | 1 | 0
| 0 | without unit
| 1 | °C
| 2 | °F
dP |  | Decimal point (max. number of digits behind the decimal point) | 0 | 0
| 0 | no digit behind the decimal point
| 1 | 1 digit behind the decimal point
| 2 | 2 digits behind the decimal point
| 3 | 3 digits behind the decimal point
LEd |  | Function allocation of status LEDs1/2/3/4 | 0 | 10
| 10 | OUT1, OUT2, OUT3, OUT4
| 11 | Heating, alarm 1, alarm 2, alarm 3
| 12 | Heating, cooling, alarm 1, alarm 2
| 13 | Cooling, heating, alarm 1, alarm 2
| 14 | Bus error
| 20 | Y1, Y2, track1, track2
| 21 | Y2, Y1, track1, track2
| 22 | track1, track2, track 3, track4
dISP | 0...10 | Display luminosity | 5 | 5
C.dEL | 0..200 | Modem delay [ms] Additional delay time, before the received message is evaluated in the Modbus. This time is required, unless messages are transferred continuously during modem transmission. | 0 | 0
dP.AD | 0...126 | Profibus address | 126
bc.uP |  | Behaviour as backup controller | 0
| 0 | No backup functionality
| 1 | With backup functionality
FrEq |  | Switching 50 Hz / 60 Hz (only visible with 8800/8840 configurator!) | 0
| 0 | 50 Hz
| 1 | 60 Hz
ICof |  | Block controller off (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Blocked
IAda |  | Block auto tuning (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Blocked
IExo |  | Block extended operating level (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Blocked
ILat |  | Suppression error storage | 0
| 0 | Released
| 1 | Blocked
Pass | OFF...9999 | Password (only visible with 8800/8840 configurator!) | OFF
IPar |  | Block parameter level (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Blocked
ICnf |  | Block configuration level (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Block
ICal |  | Block calibration level (only visible with 8800/8840 configurator!) | 0
| 0 | Released
| 1 | Blocked
CDis3 |  | Display 3 controller operating level (only visible with 8800/8840 configurator!) | 2
| 0 | No value / only text
| 1 | Display of value
| 2 | Output value as bargraph
| 3 | Control deviation as bargraph
| 4 | Process value as bargraph
TDis3 | 2...60 | Control deviation as bargraph | 10
PDis3 |  | Display 3 display alternation time [s] (only visible with 8800/8840 configurator!) | 0
| 0 | Segm.-No., Segm.-Type, Progr.-remaining time
| 1 | Segm.-No., Segm.-Type, Segm.-remaining time
| 2 | Segm.-No., Segm.-Type, net-time
| 3 | Segm.-No., Segm.-Type, Progr.-remaining time
| 4 | Segm.-No., Segm.-Type, Segm.-remaining time
| 5 | Segm.-No., Segm.-Type, net-time
T.dis3 | 8 characters | Text display 3
T.InF1 | 8 characters | Text Inf.1
T.InF2 | 8 characters | Text Inf.2
t.PrG01 | 8 characters | Text Program 1
t.PrG02 | 8 characters | Text Program 2
…
t.PrG16 | 8 characters | Text Program16
Lin (only visible with 8800/8840 configurator)
Name | Value range | Description | Default | Last settings
Lin |  | Linearization for inputs INP1 or INP3 Access to this table is always with selection special thermocouple for InP.1 or InP.3or with setting S.Lin = 1: special linearization for linearization. Default: KTY 11-6 (0...4,5 kOhm) | 1036
In.1 | -999.0..99999 | Input value 1 The signal is in [µV] or in [Ω] dependent of input type | -49,94
Ou.1 | 0,001...9999 | Output value 1 Signal assigned to In.1 | 1150
In.2 | -999.0..99999 | Input value 2 The signal is in [µV] or in [Ω] dependent of input type | -38,94
Ou.2 | 0,001...9999 | Output value 2 Signal assigned to In.2 | :
: : | : : |  | :
In.16 | -999.0..99999 | Input value 16 The signal is in [µV] or in [Ω] dependent of input type | 4470
Ou.16 | 0,001...9999 | Output value 1 6 Signal assigned to In.16 | 150,0
i Resetting the controller configuration to factory setting (Default) -> chapter 11.1 (page 82) - 8800/8840 configurator - the engineering tool for the West controller series 3 engineering tools with different functionality facilitating 8800 process controller configuration and parameter setting are available (see chapter 9: Accessory equipment with ordering information). In addition to configuration and parameter setting, 8800/8840 configurator is used for data acquisition and offers long-term storage and print functions. 8800/8840 configurator is connected to 8800 process controller via the front-panel interface by means of PC (Windows 95 / 98 / NT) and a PC adaptor. Description 8800/8840 configurator: see chapter 8: 8800/8840 configurator (page 74).
Setting parameters (2nd)
Cntr (Control and self-tuning)
Name | Value range | Description | Default | Last settings
Pb1 | 1...9999 1 | Proportional band 1 (heating) in phys. dimensions (e.g. °C) | 100 | 8
Pb2 | 1...9999 1 | Proportional band 2 (cooling) in phys. dimensions (e.g. °C) | 100
ti1 | 0,1...9999 | Integral action time 1 (heating) [s] | 180 | 15
ti2 | 0,1...9999 | Integral action time 2 (cooling) [s] | 180
td1 | 0,1...9999 | Derivative action time 1 (heating) [s] | 180 | OFF
td2 | 0,1...9999 | Derivative action time 2 (cooling) [s] | 180
t1 | 0,4...9999 | Minimal cycle time 1 (heating) [s]. The minimum impulse is 1/4x t1 | 10 | 1
t2 | 0,4...9999 | Minimal cycle time 2 (heating) [s]. The minimum impulse is 1/4x t2 | 10
SH | 0...9999 | Neutral zone or switching differential for on-off control (phys. dimensions) | 2
HYS.L | 0...9999 | Switching difference Low signaller [engineering unit] | 1
HYS.H | 0...9999 | Switching difference High signaller [engineering unit] | 1
d.SP | -1999...9999 | Trigger point seperation for additional contact ∆ /Y/ Off [phys. dimensions] | 100
tP | 0,1...9999 | Minimum impulse [s] | OFF
tt | 3...9999 | Motor travel time [s] | 60
Y2 | -100...100 | 2. correcting variable | 0 | 0
Y.Lo | -120...120 | Lower output limit [%] | 0 | 0
Y.Hi | -120...120 | Upper output limit [%] | 100 | 100
Y2 | -100...100 | 2. correcting variable | 0
Y.0 | -100...100 | Working point for the correcting variable [%] | 0 | 0
Yn.H | -100...100 | Limitation of the mean value Ym [%] | 5
L.Yn | 0...9999 | Max. deviation xw at the start of mean value calculation [phys. dimensions] | 8
E.H20 | -1999...9999 | Min. temperature for water cooling. Below the set temperature no water cooling happens | 0
t.on | 0,1...9999 | Impulse lenght for water cooling. Fixed for all values of controller output.The pause time is varied. | 1
t.oFF | 1...9999 | Min. pause time for water cooling. The max. effective controller output results from t.on/(t.on+t.off)·100% | 10
F.H20 | 0,1...9999 | Modification of the (non-linear) water cooling characteristic (see page 50) | 1
oFFS | -120...120 | Zero offset | 0
tEnp | 0...9999 | Sensor temperature (in engineering units e.g. °C) With oxygen measurement (O2) (see page 70) | 750
1 Valid for ConF/ othr/ dP = 0. With dP = 1 / 2 / 3 also 0,1 / 0,01 / 0,001 is possible.
PAr.2 (2. set of paramenters)
Name | Value range | Description | Default | Last settings
Pb12 | 1...9999 | Proportional band 1 (heating) in phys. dimensions (e.g. °C), 2. parameter set | 100 | 100
Pb22 | 1...9999 | Proportional band 2 (cooling) in phys. dimensions (e.g. °C), 2. parameter set | 100
ti22 | 0,1...9999 | Integral action time 2 (cooling) [s], 2. parameter set | 180
ti12 | 0,1...9999 | Integral action time 1 (heating) [s], 2. parameter set | 180 | 180
td12 | 0,1...9999 | Derivative action time 1 (heating) [s], 2. parameter set | 180 | 180
cntr (Control and self-tuning) | 0,1...9999 | Derivative action time 2 (cooling) [s], 2. parameter set | 180
SEtP (Set point and process-value)
Name | Value range | Description | Default | Last settings
SP.L0 | -1999...9999 | Set-point limit low for Weff (effective set-point) | -100 | 800
SP.Hi | -1999...9999 | Set-point limit high for Weff (effective set-point) | 1200 | 1200
SP.2 | -1999...9999 | Set-point 2 (2nd set point) | 0 | 0
r.SP | 0...9999 | Set-point gradient [/min] | OFF | OFF
SP | -1999...9999 | Set-point (only visible with 8800/8840 configurator!) | 0
SP.L0 and SP.Hi should be within the limits of rnGH and rnGL see configuration r Controller, page 17
InP.1
Name | Value range | Description | Default | Last settings
InL.1 | -1999...9999 | Input value for the lower scaling point | 0 | 4
OuL.1 | -1999...9999 | Displayed value for the lower scaling point | 0 | 0
InH.1 | -1999...9999 | Input value for the upper scaling point | 20 | 20
OuH.1 | -1999...9999 | Displayed value for the lower scaling point | 20 | 3000
t.F1 | 0,0...9999 | Filter time constant [s] | 0,5 | 0.5
Etc.1 | 0...100 (°C) 32...212 (°F) | External cold-junction reference temperature (external TC) | OFF
Lim (Limit value functions)
Name | Value range | Description | Default | Last settings
L.1 | -1999...9999 | Lower limit 1 | -10 | 0
H.1 | -1999...9999 | Upper limit 1 | 10 | 3000
HYS.1 | 0...9999 | Hysteresis limit 1 | 1 | 0
dEL.1 | 0...9999 | Alarm delay from limit value 1 | 0 | 5
L.2 | -1999...9999 | Lower limit 2 | OFF
H.2 | -1999...9999 | Upper limit 2 | OFF
HYS.2 | 0...9999 | Hysteresis limit 2 | 1
dEL.2 | 0...9999 | Alarm delay from limit value 2 | 0
L.3 | -1999...9999 | Lower limit 3 | OFF
H.3 | -1999...9999 | Upper limit 3 | OFF
HYS.3 | 0...9999 | Hysteresis limit 3 | 1
dEL.3 | 0...9999 | Alarm delay from limit value 3 | 0
i Resetting the controller configuration to factory setting (Default) -> chapter 11.1 (page 82)

## P6100 settings
Product Information Mode Parameters
|  |  | Default Values |  | Actual settings
Parameter | Possible Values | Upper/Main Display | Lower Display (or 1st Legend)* | 1/8 DIN Indicator Units Display | Upper/Main Display | Lower Display (or 1st Legend)*
Input type | Universal input | Un 1 | In_1 | t | Un 1 | In_1
Option 1 module type | No option fitted | nonE | Opn 1 | 1 | L in | Opn 1
| Relay | rLY
| SSR drive | SSr
| Triac | tr 1
| Linear voltage / current output | L in
Option 2 module type | No option fitted. | nonE | OPn2 | 2 | nonE | oPn2
| Relay | rLY
| Dual relay | drLy
| SSR drive | SSr
| Triac | tr 1
| Linear voltage / current output | L in
Option 3 module type | No option fitted. | nonE | Opn3 | 3 | nonE | OPn3
| Relay | drLY
| Dual relay | rLY
| SSR drive | SSr
| Linear voltage / current output | Lin
| 24V Transmitter power supply | dc24
Auxiliary option A module type | No option fitted | nonE | OPnA | A | nonE | OPnA
| RS485 comms | r485
| Digital input | diGi
| Basic Auxiliary input | rSPi
|  |  | Default Values |  | Actual settings
Parameter | Possible Values | Upper/Main Display | Lower Display (or 1st Legend)* | 1/8 DIN Indicator Units Display | Upper/Main Display | Lower Display (or 1st Legend)*
Auxiliary option B module type | No option fitted | nonE | 0Pnb | Not Applicable
| Full Auxiliary input and digital input 2 | rSPi
Firmware | Value displayed is firmware type number |  | FLJ | F | 204A | FLJ
Issue No. | Value displayed is firmware issue number |  | ISS | n | 12 | ISS
Product Rev | Value displayed is Product Revision Level. |  | PrL | r | oAoF | PrL
Level
Date of manufacture | Manufacturing date code (mmyy) |  | d0г7 | d | 1017 | d0r7
Serial number 1 | First four digits of serial number |  | Sn1 | A | 8924 | Sn1
Serial number 2 | Second four digits of serial number |  | Sn2 | b | 5701 | Sn2
Serial number 3 | Last four digits of serial number |  | Sn3 | c | 0010 | Sn3
Configuration Mode Parameters
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Input type and Range | InPt | bC | B type: 100 to 1824 ºC | JC for Europe JF for USA | Always | InPt | 4_20
|  | bF | B type: 211 to 3315 ºF
|  | CC | C type: 0 to 2320 ºC
|  | CF | C type: 32 to 4208 ºF
|  | JC | J type: -200 to 1200 ºC
|  | JF | J type: -328 to 2192 ºF
|  | J.C | J type: -128.8 to 537.7 ºC with decimal point
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Input type and Range |  | J.F | J type: -199.9 to 999.9 ºF with decimal point | JC for Europe JF for USA | Always
|  | PC | K type: -240 to 1373 ºC
|  | PF | K type: -400 to 2503 ºF
|  | P.C | K type: -128.8 to 537.7 ºC with decimal point
|  | P.F | K type: -199.9 to 999.9 ºF with decimal point
|  | LC | L type: 0 to 762 ºC
|  | LF | L type: 32 to 1403 ºF
|  | L.C | L type: 0.0 to 537.7 ºC with decimal point
|  | L.F | L type: 32.0 to 999.9 ºF with decimal point
|  | nC | N type: 0 to 1399 ºC
|  | nF | N type: 32 to 2551 ºF
|  | rC | R type: 0 to 1759 ºC
|  | rF | R type: 32 to 3198 ºF
|  | SC | S type: 0 to 1762 ºC
|  | SF | S type: 32 to 3204 ºF
|  | tC | T type: -240 to 400 ºC
|  | tF | T type: -400 to 752 ºF
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Input type and Range |  | t.C | T type: -128.8 to 400.0 ºC | JC for Europe JF for USA | Always
with decimal point
|  | t.F | T type: -199.9 to 752.0 ºF
with decimal point
|  | P24C | PtRh20% vs PtRh40%:0 to 1850 ºC
|  | P24F | PtRh20% vs PtRh40%:32 to 3362
|  | PtC | Pt100: -199 to 800 ºC
|  | PtF | Pt100: -328 to 1472 ºF
|  | Pt.C | Pt100: -128.8 to 537.7 ºC with decimal point
|  | Pt.F | Pt100: -199.9 to 999.9 ºF with decimal point
|  | 0_20 | 0 to 20mA DC
|  | 4_20 | 4 to 20mA DC
|  | 0_50 | 0 to 50mV DC
|  | 10.5 | 10 to 50mV DC
|  | 0_5 | 0 to 5V DC
|  | 1_5 | 1 to 5V DC
|  | 0_10 | 0 to 10V DC
|  | 2_10 | 2 to 10V DC
Scale Range Upper Limit | ruL | Scale Range Lower Limit +100 to Range Max |  | Linear inputs =1000 (°C/°F inputs = max range) | Always | ruL | 13
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Scale Range Lower Limit | rLL | Range Min. to Scale range Upper Limit - 100 |  | Linear = 0 (°C/°F = min range) | Always | rLL | 0
Decimal point position | dPoS | 0 | Decimal point position in non-temperature ranges. 0 = XXXX 1 = XXX.X 2 = XX.XX 3 = X.XXX | 1 | InPt = mV, V or mA | dPoS | 2
1
2
3
Control Type | CtYP | SnGL | Primary control | SnGL | Always | CtYP | SnGL
|  | duAL | Primary and Secondary control (e.g. for heat & cool)
Primary Output Control Action | CtrL | rEu | Reverse Acting | rEu | Always | CtrL | rEu
|  | d ir | Direct Acting
Alarm 1Type | ALA 1 | P_Hi | Process High Alarm | P_H1 | Always | ALA 1 | nonE
|  | P_Lo | Process Low Alarm
|  | dE | Deviation Alarm
|  | bAnd | Band Alarm
|  | nonE | No alarm
Process High Alarm 1 value* | PhA 1 | Range Min. to Range Max. Parameter repeated in Setup Mode |  | Range Max. | ALA 1 = P_Hi
Process Low Alarm 1 value* | PLA 1 | Range Min. to Range Max. Parameter repeated in Setup Mode |  | Range in | ALA 1 = P_Lo
Deviation Alarm 1 Value* | dAL 1 | ±span from setpoint Parameter repeated in Setup Mode |  | 5 | ALA 1 = dE
Band Alarm 1 Value* | bAL 1 | 1 LSD to full span from setpoint Parameter repeated in Setup Mode |  | 5 | ALA 1 = bAnd
Alarm 1 Hysteresis* | AHY 1 | 1 LSD to 100% of span (in display units) on “safe” side of alarm point Parameter repeated in Setup Mode |  | 1 | Always
Alarm 2 Type | ALA2 | Same options as Alarm 1 type |  | P_Lo | Always | ALA 2 | nonE
Process High Alarm 2 value* | PhA2 | Range Min. to Range Max. Parameter repeated in Setup Mode |  | Range Max. | ALA2 = P_Hi
Process Low Alarm 2 value* | PLA2 | Range Min. to Range Max. Parameter repeated in Setup Mode |  | Range Min | ALA2 = P_Lo
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Deviation Alarm 2 value* | dAL2 | ±span from setpoint Parameter repeated in Setup Mode |  | 5 | ALA2 = dE
Band Alarm 2 value* | bAL2 | 1 LSD to full span from setpoint .Parameter repeated in Setup Mode |  | 5 | ALA2 = bAnd
Alarm 2 Hysteresis* | AHY2 | 1 LSD to 100% of span (in display units) on “safe” side of alarm point.Parameter repeated in Setup Mode |  | 1 | Always
Loop Alarm Enable | LAEn | diSA (disabled) or EnAb (enabled) |  | diSA | Always | LAEn | diSA
Alarm Inhibit | Inh 1 | nonE | No alarms Inhibited | nonE | Always | Inh 1 | nonE
|  | ALA 1 | Alarm 1 inhibited
|  | ALA2 | Alarm 2 inhibited
|  | both | Alarm 1 and alarm 2 inhibited
Output 1 Usage | USE 1 | Pr 1 | Primary Power | Pr 1 | Opn 1 is not nonE | USE 1 | rEtS
|  | Sec | Secondary Power
|  | A1_d | Alarm 1, Direct Acting
|  | A1_r | Alarm 1, Reverse Acting
|  | A2_d | Alarm 2, Direct Acting
|  | A2_r | Alarm 2, Reverse Acting
|  | LP_d | Loop Alarm, Direct Acting
|  | LP_r | Loop Alarm, Reverse Acting
|  | Or_d | Logical Alarm 1 OR Alarm 2 Direct Acting
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Output 1 Usage | USE 1 | Or_r | Logical Alarm 1 OR Alarm 2 Reverse Acting | Pr 1 | Opn 1 is not nonE
|  | Ar_d | Logical Alarm 1 AND Alarm 2, Direct Acting
|  | Ar_r | Logical Alarm 1 AND Alarm 2, Reverse Acting
|  | rEtS | Retransmit SP Output
|  | rEtP | Retransmit PV Output
Linear Output 1 Range | tYP 1 | 0 to 5 V DC output 1 |  | 0_10 | 0Pn 1 = L in | tYP 1 | 4_20
0 to 10 V DC output
2 to 10 V DC output
0 to 20 mA DC output
4 to 20 mA DC output
Retransmit Output 1 Scale maximum | ro IH | -1999 to 9999 Display value at which output will be maximum |  | Range max | USE 1 = rEtS or rEtP | ro IH | 13
Retransmit Output 1 Scale minimum | ro IL | -1999 to 9999 Display value at which output will be minimum |  | Range min | USE 1 = rEtS or rEtP | ro IL | 0
Output 2 Usage | USE2 | Same options as Output 1. |  | SEc if dual control selected else A2_d | OPn2 is not nonE
Linear Output 2 Range | TYP2 | Same options as Output 1. |  | 0_ 10 | OPn2 = L in
Retransmit Output 2 Scale maximum | ro2H | -1999 to 9999 Display value at which output will be maximum |  | Range max | USE2 = rEtS or rEtP
Retransmit Output 2 Scale minimum | ro2L | -1999 to 9999 Display value at which output will be minimum |  | Range min | USE2 = rEtS or rEtP
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Output 3 Usage** | USE3 | Same options as Output 1. |  | A 1_d (rEtP if linear) | OPn3 is not nonE
Linear Output 3 Range | tYP3 | Same options as Output 1. |  | 0_ 10 | OPn3 = L in
Retransmit Output 3 Scale maximum | ro3H | -1999 to 9999 Display value at which output will be minimum |  | Range max | USE3 = rEtS or rEtP
Retransmit Output 3 Scale minimum | ro3L | -1999 to 9999 Display value at which output will be minimum |  | Range min | USE3 = rEtS or rEtP
Display Strategy | diSP | 1, 2, 3, 4, 5, 6 or 7 (see Operator Mode) |  | 1 | Always | diSP | 2
Communications Protocol | Prot | ASC 1 | ASCII | r7bn | OPnA = r4B5
|  | r7bn | Modbus with no parity
|  | r7bE | Modbus with Even Parity
|  | r7bo | Modbus with Odd Parity
Bit rate | bAud | 1.2 | 1.2 kbps | 4.8 | 0PnA = r4B5
|  | 2,4 | 2.4 kbps
|  | 4.8 | 4.8 kbps
|  | 9.6 | 9.6 kbps
|  | 19.2 | 19.2 kbps
Communications Address | Addr | 1 | Unique address assigned to the instrument in the range of 1 to 255 | 1 | OPnA = r485
Communicatios Write Enable | CoEn | r_o | Read only. Comms writes ignored | r_ LJ | Always
|  | r_LJ | Read / Write. Writing via Comms is possible
Digital Input 1 Usage | diG1 | diS1 | Setpoint 1 / Setpoint 2 Select** | diS1 | OPnA = diGi
Parameter | Lower Display | Upper Display | Description | Default Settings |  | Actual settings
|  |  |  | Default Value | When Visible | Lower Display | Upper Display
Digital Input 1 Usage |  | diAS | Automatic / Manual Select** | diS1 | OPnA = diGi
Digital Input 2 Usage | diG2 | diS1 | Setpoint 1 / Setpoint 2 Select** | dirS | 0Pnb = rSP1
|  | diAS | Automatic / Manual Select**
|  | dirS | Remote / Local Setpoint Select
Remote Setpoimt Input Range | r inP | 0_20 | 0 to 20mA DC input | 0_10 | OPnA or Opnb = rSP i
|  | 4_20 | 4 to 20mA DC input
|  | 0_10 | 0 to 10V DC input
|  | 2_10 | 2 to 10V DC input
|  | 0_5 | 0 to 5V DC input
|  | 1_5 | 1 to 5V DC input
|  | 100 | 0 to 100mV DC input |  | 0Pnb = rSP i
|  | Pot | Potentiometer (≥2KΩ)
Remote Setpoint Upper Limit | rSPu | -1999 to 9999 RSP value to be used when RSP input is at maximum |  | Range max | 0PnA = rSP i
Remote Setpoint Lower Limit | rSPL | -1999 to 9999 RSP value to be used when RSP input is at minimum |  | Range min | 0PnA = rSP i
Remote Setpoint Offset | rSPo | Offset applied to RSP value. Constrained within Scale Range Upper Limit and Scale Range Lower Limit. |  | 0 | 0PnA = rSP i
Configuration Mode Lock Code | Cloc | 0 to 9999 |  | 20 | Always | Cloc | 20
Set Up Mode Parameters
Parameter | Lower Display | Upper Display Adjustment Range | Default Value | When Visible | Actual settings
[5 empty] | Lower Display | Upper Display
Input Filter Time constant | FiLt | OFF, 0.5 to 100.0 secs in 0.5 sec increments | 2 | Always | FiLt | 2
Parameter | Lower Display | Upper Display Adjustment Range | Default Value | When Visible | Actual settings
[5 empty] | Lower Display | Upper Display
Process Variable Offset | OFFS | ±Span of controller | 0 | Always | OFFS | 0
Parameter | Lower Display | Upper Display Adjustment Range | Default Value | When Visible | Actual settings
[5 empty] | Lower Display | Upper Display
Primary Power | PPLJ | The current Primary Output Power. Read Only. | N/A | Always | PPLJ | 82
Secondary Power | SPLJ | The current Secondary Output power. Read Only. | N/A | CtYP = duAL
Primary Output Proportional Band | Pb_P | 0.0% (ON/OFF control) and 0.5% to 999.9% of input span. | 10 | Always | Pb_P | 691.9
Secondary Output Proportional Band | Pb_S | 0.0% (ON/OFF control) and 0.5% to 999.9% of input span. | 10 | CtYP = duAL
Automatic Reset (Integral Time Constant) | ArSt | 0.01 to 99.59 (1 sec to 99 mins 59 secs) and OFF (blank) | 5 | Pb_P is not 0.0 | ArSt | 0.01
Rate (Derivative Time Constant) | RAtE | 0.00 to 99.59 (OFF to 99 mins 59 secs) | 1.15 | Pb_P is not 0.0 | RAtE | 0
Overlap/Deadband | OL | -20% to +20% of the sum of the Primary and Secondary Proportional Bands | 0 | Pb_P is not 0.0
Manual Reset (Bias) | biAS | 0% to 100% (-100% to 100% if (CtYP = duAL) | 25 | Pb_P is not 0.0 | biAS | 25
Primary Output ON/OFF Differential | diFP | 0.1% to 10.0% of input span (enter in % span) | 0.5 | Pb_P = 0.0
Secondary Output ON/OFF Differential | diFS | 0.1% to 10.0% of input span (enter in % span) | 0.5 | Pb_P = 0.0
Primary and Secondary Output ON/OFF Differential | diFF | 0.1% to 10.0% of input span (enter in % span) | 0.5 | Pb_P and Pb_s = 0.0
Setpoint Upper Limit | SPuL | Current Setpoint value to Scale Range Maximum | Range Max. | Always | SPuL | 13
Parameter | Lower Display | Upper Display Adjustment Range | Default Value | When Visible | Actual settings
[5 empty] | Lower Display | Upper Display
Setpoint Lower Limit | SPLL | Scale Range Minimum to current Setpoint value | Range Min | Always | SPLL | 0
Primary (Heat) Output Upper Power Limit | OPuL | 0% to 100% of full power | 100 | Pb_P is not 0.0 | OPuL | 100
Output 1 Cycle Time | Ct 1 | 0.5, 1, 2, 4, 8, 16, 32, 64, 128, 256 or 512 secs. Not applicable to linear outputs | 32 | USE 1 = Pr 1 or SEc or buS
Output 2 Cycle Time | Ct2 | 0.5, 1, 2, 4, 8, 16, 32, 64, 128, 256 or 512 secs. Not applicable to linear outputs | 32 | USE 2 = Pr 1 or SEc or buS
Output 3 Cycle Time | Ct3 | 0.5, 1, 2, 4, 8, 16, 32, 64, 128, 256 or 512 secs. Not applicable to linear outputs | 32 | USE 3 = Pr 1 or SEc or buS
Process High Alarm 1 value* | PhA 1 | Range Min. to Range Max. | Range Max. | ALA 1 = P_Hi
Process Low Alarm 1 value* | PLA 1 | Range Min. to Range Max. | Range Min. | ALA 1 = P_Lo
Deviation Alarm 1 value* | dAL 1 | ± span from setpoint | 5 | ALA 1 = dE
Band Alarm 1 value* | bAL 1 | 1 LSD to full span from setpoint. | 5 | ALA 1 = bAnd
Alarm 1 Hysteresis* | AHY 1 | Up to 100% of span | 1 | Always
Process High Alarm 2 value* | PhA 2 | Range Min. to Range Max. | Range Max. | ALA 2 = P_Hi
Process Low Alarm 2 value* | PLA 2 | Range Min. to Range Max. | Range Min | ALA 2 = P_Lo
Deviation Alarm 1 value* | dAL 2 | ± span from setpoint | 5 | ALA 2 = dE
Band Alarm 2 value* | bAL 2 | 1 LSD to full span from setpoint. | 5 | ALA 2 = bAnd
Alarm 2 Hysteresis* | AHY 2 | Up to 100% of span | 1 | Always
Loop Alarm Time* | Lat 1 | 1 sec to 99 mins. 59secs. Only applies if primary proportional band = 0 | 99.59 | LAEn = EnAb
Auto Pre-tune enable/disable | Apt | diSA disabled or EnAb enabled | diSA | Always | Apt | diSA
Parameter | Lower Display | Upper Display Adjustment Range | Default Value | When Visible | Actual settings
[5 empty] | Lower Display | Upper Display
Manual Control select enable / disable | PoEn | diSA disabled or EnAb enabled | diSA | Always | PoEn | diSA
Setpoint Select shown in Operator Mode, enable / disable | SSEn | diSA disabled or EnAb eenabled | diSA | Slot A or B fitted with RSP module
Setpoint ramp shown in operator mode, enable / disable | SPr | diSA disabled or EnAb enabled | diSA | Always | SPr | diSA
SP Ramp Rate Value | rP | 1 to 9999 units/hour or OFF (blank) | Blank | Always | rP | Blank
Setpoint Value | SP | Within scale range upper and lower limits | Range minimum | Always | SP | 5
Local Setpoint Value | LSP _LSP or ≡LSP | Within scale range upper and lower limits _ or ≡ before the legend indicates if this is the currently active SP | Range minimum. | 0PnA or 0Pnb = rSP 1
Setpoint 1 Value | SP1 _SP1 or ≡SP1 | Within scale range upper and lower limits _ or ≡ before the legend indicates if this is the currently active SP | Range minimum. | diG 1 or d iG2 = d iS1
Setpoint2 Value | SP2 _SP2 or ≡SP2 | Within scale range upper and lower limits _ or ≡ before the legend indicates if this is the currently active SP | Range minimum. | diG 1 or d iG2 = d iS2
Set-up Lock Code | SLoc | 0 to 9999 | 10 | Always | SLoc | 10